## Sheet1
FPix to DTB adapter v2.1 test protocol
S/N | 14
1. Tests for shorts
Material: | Multimeter in resistor mode
Instructions: | Test all connections listed in the table below and record the values observed. Use the test points on the adapter board.
A flex cables needs to be inserted, side with small pads up. ∞ means maximal reading of Ohmmeter
Point 1 | Point 2 | Expected (Ω) | Tolerance (±Ω) | Observed (Ω) | Ok?
GND | VA | ∞ | N/A | ∞ | ok | 0
GND | VD | ∞ | N/A | ∞ | ok | 0
GND | 3.3V | 20 | 0.02828 | 6.67 | NOK | 1
GND | RTD | 10 | 0.01 | 6.67 | NOK | 1
GND | OutA+ | ∞ | N/A | ∞ | ok | 0
GND | OutA- | ∞ | N/A | ∞ | ok | 0
GND | RDa+ | ∞ | N/A | ∞ | ok | 0
GND | RDa- | ∞ | N/A | ∞ | ok | 0
GND | SD+ | ∞ | N/A | ∞ | ok | 0
GND | SD- | ∞ | N/A | ∞ | ok | 0
GND | Trig+ | ∞ | N/A | ∞ | ok | 0
GND | Trig- | ∞ | N/A | ∞ | ok | 0
GND | Clk+ | ∞ | N/A | ∞ | ok | 0
GND | Clk- | ∞ | N/A | ∞ | ok | 0
OutA+ | OutA- | 120 | 6 | 120.3 | ok | 0
OutA- | RDa+ | ∞ | N/A | 0.1 | NOK | 1
RDa+ | RDa- | 120 | 6 | 120.4 | ok | 0
RDa- | SD+ | ∞ | N/A | ∞ | ok | 0
SD+ | SD- | 91 | 4.55 | 90.6 | ok | 0
SD- | Trig+ | ∞ | N/A | ∞ | ok | 0
Trig+ | Trig- | 91 | 4.55 | 91.1 | ok | 0
Trig- | Clk+ | ∞ | N/A | ∞ | ok | 0
Clk+ | Clk- | 91 | 4.55 | 91.2 | ok | 0
SDATA1+ | SDATA1- | 120 | 6 | 120.4 | ok | 0
TOUT+ | TOUT- | 120 | 6 | 120.4 | ok | 0
SDA+ | SDA- | 169 | 5.3205 | 168.5 | ok | 0
CTR+ | CTR- | 169 | 5.3205 | 169 | ok | 0
CLK+ | CLK- | 169 | 5.3205 | ∞ | #VALUE! | #VALUE!
RTD | 3.3V | 30 | 0.051962 | 0.2 | NOK | 1
VA | OutA- | ∞ | N/A | 0.1 | NOK | 1
VA | RDa+ | ∞ | N/A | 0.1 | NOK | 1
VD | Clk- | ∞ | N/A | 0.1 | NOK | 1
VD | 3.3V | ∞ | N/A | ∞ | ok | 0
HV (on cable) | GND | ∞ | N/A | ∞ | ok | 0
Overall result: | #VALUE!
2. LED test
Material: | DTB, connected to psi46test or pXar
Instructions: | Connect adapter (still with flex cable) to DTB. Issue power on command. LED should light up.
Result:
3. HV test
Material: | DTB, connected to psi46test or pXar, hooked up to Keithley 2410
Instructions: | Connect adapter (still with flex cable) to DTB. Turn HV on.
Set Keithley voltage to -1000 V, current compliance 100 µA. Turn voltage on.
Keithley should not trip (ok within the first second)
Current rading after 20 s: |  |  | µA
Limit: |  | 0.01 | µA
Result:
4. Summary
Module test:
Tester: | Frank Meier |  |  | Test date: | 2014-07-09