## Sheet1
FPix to DTB adapter v2.1 test report
S/N of adapter card: |  | 67
Test equipment: |  | Brand | S/N
Multimeter |  | Fluke | 79980109 |  | Resistance if shorted: |  | 0.2 | Ω
DTB |  | PSI
HV supply |  | Keithley
1. Tests for shorts and connections
Material: | Multimeter in resistor mode
Instructions: | Test all connections listed in the table below and record the values observed. Use the test points on the adapter board.
A flex cables needs to be inserted, side with small pads up. ∞ means maximal reading of Ohmmeter
Disregard, structure to get overall PASS
Point 1 | Point 2 | Unit | Expected | ± Tolerance | Observed | Ok? |  | Cable | Point | Observed (Ω) | Ok? |  | Outside printing range
GND | VA | Ω | ∞ | N/A | ∞ | ok |  | 1 | OutA+ |  | NOK |  | 0 | 1
GND | VD | Ω | ∞ | N/A | ∞ | ok |  | 2 | OutA- |  | NOK |  | 0 | 1
GND | 3.3V | kΩ | 20 | 0.02828 | 19.99 | ok |  | 3 | VA |  | NOK |  | 0 | 1
GND | RTD | kΩ | 10 | 0.01 | 10 | ok |  | 4 | VA |  | NOK |  | 0 | 1
GND | OutA+ | Ω | ∞ | N/A | ∞ | ok |  | 5 | RDa+ |  | NOK |  | 0 | 1
GND | OutA- | Ω | ∞ | N/A | ∞ | ok |  | 6 | RDa- |  | NOK |  | 0 | 1
GND | RDa+ | Ω | ∞ | N/A | ∞ | ok |  | 7 | SD+ |  | NOK |  | 0 | 1
GND | RDa- | Ω | ∞ | N/A | ∞ | ok |  | 8 | SD- |  | NOK |  | 0 | 1
GND | SD+ | Ω | ∞ | N/A | ∞ | ok |  | 13 | Trig+ |  | NOK |  | 0 | 1
GND | SD- | Ω | ∞ | N/A | ∞ | ok |  | 14 | Trig- |  | NOK |  | 0 | 1
GND | Trig+ | Ω | ∞ | N/A | ∞ | ok |  | 15 | Clk+ |  | NOK |  | 0 | 1
GND | Trig- | Ω | ∞ | N/A | ∞ | ok |  | 16 | Clk- |  | NOK |  | 0 | 1
GND | Clk+ | Ω | ∞ | N/A | ∞ | ok |  | 17 | VD |  | NOK |  | 0 | 1
GND | Clk- | Ω | ∞ | N/A | ∞ | ok |  | 18 | VD |  | NOK |  | 0 | 1
OutA+ | OutA- | Ω | 120 | 6 | 119.8 | ok |  | 19 | 3.3V |  | NOK |  | 0 | 1
OutA- | RDa+ | Ω | ∞ | N/A | ∞ | ok |  | 20 | RTD |  | NOK |  | 0 | 1
RDa+ | RDa- | Ω | 120 | 6 | 120.3 | ok |  | Limit |  | 1 |  |  | 0
RDa- | SD+ | Ω | ∞ | N/A | ∞ | ok | [6 empty] | 0
SD+ | SD- | Ω | 91 | 4.55 | 90.9 | ok |  | Counting from top to bottom |  |  |  |  | 0
SD- | Trig+ | Ω | ∞ | N/A | ∞ | ok |  | No. 9..12 are the HV plus surround empty pads |  |  |  |  | 0
Trig+ | Trig- | Ω | 91 | 4.55 | 90.8 | ok | [6 empty] | 0
Trig- | Clk+ | Ω | ∞ | N/A | ∞ | ok |  | Cable | Point | Observed (Ω) | Ok? |  | 0
Clk+ | Clk- | Ω | 91 | 4.55 | 91.1 | ok |  | 1 | GND |  | NOK |  | 0 | 1
SDATA1+ | SDATA1- | Ω | 120 | 6 | 119.8 | ok |  | 2 | GND |  | NOK |  | 0 | 1
TOUT+ | TOUT- | Ω | 120 | 6 | 120.3 | ok |  | 3 | VA |  | NOK |  | 0 | 1
SDA+ | SDA- | Ω | 169 | 5.3205 | 168.5 | ok |  | 4 | GND |  | NOK |  | 0 | 1
CTR+ | CTR- | Ω | 169 | 5.3205 | 168.8 | ok |  | 6 | GND |  | NOK |  | 0 | 1
CLK+ | CLK- | Ω | 169 | 5.3205 | 168.9 | ok |  | 7 | VD |  | NOK |  | 0 | 1
RTD | 3.3V | kΩ | 30 | 0.051962 | 29.98 | ok |  | 8 | GND |  | NOK |  | 0 | 1
VA | OutA- | Ω | ∞ | N/A | ∞ | ok |  | 9 | GND |  | NOK |  | 0 | 1
VA | RDa+ | Ω | ∞ | N/A | ∞ | ok |  | Limit |  | 0.5 |  |  | 0
VD | Clk- | Ω | ∞ | N/A | ∞ | ok | [6 empty] | 0
VD | 3.3V | Ω | ∞ | N/A | ∞ | ok |  | On back side, enumeration follows pad from |  |  |  |  | 0
HV (on cable) | GND | Ω | ∞ | N/A | ∞ | ok |  | top to bottom. Top = top on connector |  |  |  |  | 0
Overall result: | FAIL
2. LED test | [7 empty] | 4. Other tests
Material: | DTB, connected to psi46test or pXar | [6 empty] | 1) Jumper for LED installed |  |  |  | n
Instructions: | Connect adapter (still with flex cable) to DTB. Issue power on command. LED should light up. | [6 empty] | 2) Kapton foil present on back side |  |  |  | n
[8 empty] | 3) S/N sticker attached |  |  |  | n
Result:
3. HV test
Material: | DTB, connected to psi46test or pXar, hooked up to Keithley 2410 | [6 empty] | Set Keithley voltage to -100 V, current compliance 100 µA.
Instructions: | Set Keithley voltage to -1000 V, current compliance 100 µA. Turn voltage on. | [6 empty] | Turn voltage on. Measure value on flex cable end
| 1) Do not connect the adapter. Turn HV on, read after 60 sec | [6 empty] | Voltage |  |  | V | (on Multimeter)
| 2) Connect the adapter. Turn HV on, read after 60 sec | [6 empty] | Current |  |  | µA | (on Keithley)
Keithley should not trip any time.
[8 empty] | Voltage drop |  | 0 | V
| No adapter |  | µA |  |  |  |  | NB: DTB has 102 kΩ internally in series
| With adapt. |  | µA
| Difference | 0 | µA | Limit: | 0.01 | µA |  | Corrected |  | 0 | V
Result: | FAIL
Final decision: |  | REJECTED
Tester: |  |  |  | Test date: